## 1_Noise_Ranked
Rank | Noise type | How it occurs in your setup | How it corrupts the signal | Frequency band | Amplitude / severity | Features it obfuscates | How hard to remove | AD8232 vs ADS1293 | Prevalence | Ref IDs
1 | Baseline wander (respiration and posture drift) | Unavoidable even when the patient sits perfectly still. Breathing moves the chest wall, which stretches the skin under the electrode and changes the electrode-skin geometry beat to beat. Slow postural settling in the chair adds a second, slower drift. Both happen in every single recording you will ever take. | The whole trace slowly rides up and down. The critical damage is that it MOVES THE ZERO LINE. Every amplitude and every area in your feature list is measured as a distance from the isoelectric baseline, so if the baseline is wrong, all of them are wrong at once, even though the waveform shapes still look perfectly normal. | 0 to 0.5 Hz. Seated breathing sits near 0.2 to 0.35 Hz. | Can exceed the QRS amplitude. This is one of the few noises that is routinely LARGER than the signal you want. | ST-T change (worst hit of all), P-wave area, P-wave amplitude, PTFV1, T-wave amplitude, SV1, RV5, Sokolow-Lyon, Cornell voltage, QRS axis, T-wave axis, and therefore frontal QRS-T angle, P-wave axis, low QRS voltage, U wave. | MEDIUM, and there is a trap. A high-pass filter removes it, but the AHA standard sets the low-frequency cutoff at 0.05 Hz precisely to avoid distorting the ST segment (0.67 Hz is allowed only for a linear-phase digital filter). Filtering harder than that creates artificial ST depression, which is a fake version of the very feature you are trying to measure. | AD8232: its two-pole high-pass filter is aimed at heart-rate monitoring and is usually configured well above 0.05 Hz, so it will suppress wander AND distort your ST segment at the same time. ADS1293: digital filtering is programmable, so you can keep a diagnostic-grade low cutoff. | 100 percent of recordings | N1, N2, N5, N6, N7, N17
2 | Electrode motion artefact | The patient shifts in the chair, reaches for something, adjusts a sleeve, or the electrode simply peels slightly. Any relative movement between the electrode and the skin does it. This is exactly the 'small movements, electrode movements' case you described. | Three mechanisms act together: (a) the contact impedance changes, (b) skin stretch mechanically disturbs the charge double layer at the electrode-electrolyte interface, which temporarily shifts the half-cell potential, and (c) friction adds charge. The result is a large, abrupt, step-like or spike-like excursion. PhysioNet states plainly that electrode motion is generally the most troublesome noise, BECAUSE IT CAN MIMIC THE APPEARANCE OF ECTOPIC BEATS and cannot be removed by simple filters. | Mostly below 7 to 10 Hz, so it sits directly on top of the P wave, the T wave and part of the QRS. | Large. Often comparable to or larger than the QRS. | Everything, but the dangerous ones are: R-peak detection (creates FALSE BEATS, which then poison the whole RR series), fragmented QRS (fake notching), P-wave onset and offset, ST level, and any HRV measure downstream of a false beat. | HARD. This is the one the literature repeatedly names as the hardest to remove. It overlaps the signal in both frequency and shape, so no linear filter separates it. This should be the main target of your denoiser. | Both chips suffer. AD8232 has a fast restore circuit that pulls the amplifier back from saturation, but its feedback recovery switch is on for 110 ms, so roughly a tenth of a second of trace is unusable after each big event. | Very high in your setup | N1, N2, N4, N8, N9
3 | Muscle artefact (surface EMG) | Sitting upright is not passive. Postural muscles in the shoulders, chest and back stay active, and any tension, talking, shivering, cold, or holding a position causes more. Even a relaxed seated subject produces measurable EMG. | A high-frequency fuzz laid over the entire trace. The problem is spectral: surface EMG occupies roughly 5 to 500 Hz with dominant energy at 50 to 150 Hz, and the diagnostic ECG band is 0.05 to 150 Hz. THEY OVERLAP ALMOST COMPLETELY, and the overlap is worst exactly where QRS energy lives. You cannot low-pass this away without destroying the QRS. | 5 to 500 Hz, dominant 50 to 150 Hz. Overlaps QRS energy. | Moderate, but continuous rather than occasional. | QRS onset and offset (so QRS duration), P-wave onset and offset (the P wave is small enough to be buried), T-wave end (the tangent method needs a clean downstroke), QT, QT dispersion, TpTe, fragmented QRS, VAT, T-wave alternans, U wave. | HARD. Because the spectra overlap, any filter aggressive enough to remove EMG also blunts the QRS and smears the wave boundaries you need. | Same for both chips. This is a biology problem, not a chip problem. Better electrode siting and a relaxed, supported posture help more than any filter. | High in your setup | N1, N2, N3, N4, N15
4 | Powerline interference at 50 Hz | Capacitive coupling from mains wiring in the walls, phone chargers, laptop power supplies and fluorescent lighting. It gets much worse when electrode impedance is high (poor skin prep) or unbalanced between electrodes, and when the board is USB-powered from a mains adapter rather than a battery. Bangladesh mains is 50 Hz, so expect 50 Hz plus harmonics at 100 and 150 Hz, NOT the 60 Hz used in most American papers. | A steady sine wave superimposed on the ECG. It does not move the baseline and it averages out over a whole beat, so it damages fine detail rather than gross measurements. | Narrow band at 50 Hz, plus harmonics at 100 and 150 Hz. | Varies hugely with setup. Can range from invisible to larger than the P wave. | Fragmented QRS (50 Hz ripple looks exactly like the notching that defines fQRS - a serious false-positive risk), T-wave alternans, U wave, precise P-wave onset and offset, and peak-picking jitter generally. | EASY, and this is why it is ranked below the three above despite being extremely common. A notch filter removes it. But note that a sharp notch rings around the QRS, which can itself create fake notching. A well-designed notch or an adaptive canceller is worth the effort. | AD8232 CMRR is 80 dB minimum, 86 dB typical from DC to 60 Hz, and it includes a right-leg-drive amplifier to actively cancel mains. ADS1293 also has right-leg drive. Real-world rejection is far worse than the datasheet if your two electrode impedances differ, because the mismatch converts common-mode into differential. | Very high in a home setting | N3, N4, N7, N9, N10
5 | Electrode-skin impedance drift (gel drying and sweat) | Over the minutes to hours of a recording the conductive gel dries out and the adhesive weakens, while sweat forms a thin, salty, time-varying conductive layer. Both change the electrical contact continuously. | A slow, progressive change in signal amplitude and a wandering DC level. Unlike baseline wander this does not oscillate, it TRENDS, so it silently rescales your measurements over the session. Published measurements show roughly 30 percent signal degradation after 16 hours and roughly 35 percent loss of R-peak amplitude after 24 hours. | DC to very low frequency. Effectively a trend, not an oscillation. | Gradual, but cumulative and large. | All amplitude features: SV1, RV5, Sokolow-Lyon, Cornell voltage, P-wave amplitude, P-wave area, PTFV1, T-wave amplitude. USEFUL NOTE: ratio features such as Tp-e/QT and iCEB are largely immune, because the scaling cancels in the ratio. Interval features are also largely immune. | MEDIUM. Filtering cannot restore lost amplitude. The fix is skin preparation, fresh electrodes, and re-normalising amplitudes within each recording rather than across recordings. | Both chips equally affected. This is an electrode problem, not a chip problem. | High for any recording longer than a few minutes | N8, N11, N12, N13
6 | DC electrode offset and amplifier saturation | Every electrode-skin junction has its own half-cell potential, which can be hundreds of millivolts. If the two electrodes are not identical, or the gel differs, or one is disturbed, the difference appears as a large DC offset at the amplifier input. | The amplifier clips. The AD8232 accepts only plus or minus 300 mV of DC differential input, and half-cell potentials can exceed that. When it clips you get a flat, railed segment where the ECG should be, followed by a recovery transient. This is not corruption of the waveform, it is TOTAL LOSS of that stretch of data. | DC and near-DC. | Catastrophic while it lasts, but intermittent. | Every feature inside the affected window is destroyed, not just degraded. If it fires across a T wave you lose QT for that beat. If it removes beats, the RR series gets a gap and every HRV measure is affected. | NOT REMOVABLE by filtering, because the information was never recorded. The only fix is prevention: matched electrodes, good skin prep, and AC coupling. | AD8232: DC differential input range is only plus or minus 300 mV, and its fast restore circuit exists specifically for this, with a 110 ms feedback recovery switch time. ADS1293: 24-bit resolution gives far more headroom, so it tolerates offset much better. | Medium, but very damaging when it happens | N9, N10
7 | Quantisation noise (ADC resolution) | This one is your design choice, not a property of the body. The AD8232 outputs an ANALOG signal, so the resolution depends entirely on whatever ADC you feed it into. A common setup is an Arduino Uno with a 10-bit ADC. | The signal is rounded to discrete steps, so small features get flattened into staircases. Worked example, and please check it against your own gain setting: a 10-bit ADC over a 3.3 V range gives 3.22 mV per step. At a total gain of 100 that is about 32 microvolts per step at the electrodes; at a gain of 1000 it is about 3.2 microvolts per step. A P wave is roughly 100 to 250 microvolts tall, so at gain 100 the entire P wave spans only about 3 to 8 steps. | Not a frequency band. It is a resolution floor. | Depends entirely on your gain and ADC bit depth. | P-wave area (your top C1 feature - it is built from many small samples, so staircase errors accumulate), T-wave alternans (measured in single microvolts, so it can fall entirely below one step), U wave, fragmented QRS notching, fine amplitude measurements. | NOT REMOVABLE. Information destroyed at the ADC is gone. Fix it in hardware: raise the gain, use a higher-resolution ADC, or use the ADS1293. | THIS IS THE BIGGEST PRACTICAL DIFFERENCE BETWEEN YOUR TWO CHIPS. AD8232 is analog-out and inherits whatever ADC you attach. ADS1293 has integrated 24-bit converters, so quantisation stops being a limiting factor. | High if you pair AD8232 with a low-resolution ADC | N9, N10
8 | Cable and lead motion (triboelectric noise) | Wires swinging freely, rubbing against clothing, or being tugged. In a casual home setup the cables are rarely secured. | Friction and bending of the cable insulation generate and separate charge, which appears as a low-frequency artefact very similar in shape to electrode motion. It is often mistaken for electrode motion, but the fix is different. | Low frequency, similar to electrode motion. | Moderate. Usually smaller than true electrode motion. | Same set as electrode motion: R-peak detection, P-wave boundaries, ST level, fragmented QRS. | HARD to filter, EASY to prevent. Tape or clip the cables so they cannot swing. This is the cheapest single improvement in the whole list. | Both chips equally affected. It happens before the chip. | Medium to high in a casual setup | N8
9 | Amplifier and thermal noise (the hardware noise floor) | Always present. It is generated inside the front-end chip itself and sets the absolute limit on the smallest thing you can measure. | A fine random hiss. It does not distort shapes, it simply hides anything smaller than itself. For the AD8232 the datasheet gives 100 nV per root Hz spectral noise density at 1 kHz, 12 microvolts peak-to-peak over 0.1 to 10 Hz, and 14 microvolts peak-to-peak over 0.5 to 40 Hz. Put that in context: a P wave is roughly 100 to 250 microvolts, so the chip's own noise is already about 6 to 14 percent of your most important C1 feature. | Broadband. | About 14 microvolts peak-to-peak in the 0.5 to 40 Hz band for AD8232. | T-wave alternans (measured in single microvolts, so it is BELOW this floor and cannot be measured with an AD8232 at all), U wave, P-wave area precision, fine fragmented-QRS detail. | NOT REMOVABLE by filtering, but it IS reducible by averaging. Building a median beat from N beats reduces random noise by roughly the square root of N, which is why the median-beat step in your features workbook matters so much here. | AD8232 figures given above are from its datasheet. ADS1293 is a lower-noise, higher-resolution part. | 100 percent, but usually the smallest term | N9, N10
10 | Residual common-mode interference and RF pickup | Whatever mains and radio-frequency interference the right-leg drive and the amplifier's common-mode rejection fail to cancel. Phones, WiFi routers, switching power supplies and unshielded USB cables all contribute. | Broadband or narrowband interference riding on the trace. Behaves like powerline noise but at less predictable frequencies, so a fixed notch filter does not catch it. | Wide and unpredictable, often above 100 Hz. | Usually small, but occasionally severe with a bad power supply. | Fine-detail features: T-wave alternans, U wave, fragmented QRS, precise wave boundaries. | MEDIUM. Low-pass filtering helps for the high-frequency part. Prevention matters more: run on battery, keep the phone away, twist the electrode leads together. | AD8232 CMRR 80 dB min / 86 dB typ. Both chips provide right-leg drive. Rejection collapses if electrode impedances are mismatched. | Medium | N4, N9, N10
11 | Intermittent lead-off and loose contact | An electrode partly lifts, or a connector works loose. Very common with casual self-application and no clinical supervision. | Sudden dropouts, rail-to-rail swings, or long stretches of pure noise where there is no ECG at all. Unlike the noises above, this produces segments that contain NO signal, so they must be detected and discarded rather than cleaned. | Not applicable. It is a fault condition. | Total loss while it lasts. | All features in the affected window. Critically, if you do not DETECT these segments, your feature extractor will happily output confident nonsense values for them. | MUST BE DETECTED AND EXCLUDED, never cleaned. This is why the signal-quality row in your features workbook matters. | AD8232 has AC lead-off detection using a 100 kHz, 200 nA peak-to-peak excitation with a 10 to 20 megaohm impedance threshold, and a 110 microsecond detection delay. ADS1293 has both AC and DC lead-off detection. USE THIS OUTPUT - it is free ground truth for your signal-quality labels. | Medium, and higher for unsupervised home use | N9, N10
12 | Respiration-induced amplitude modulation | Breathing does not only shift the baseline. It also rotates the heart slightly and changes chest impedance, which modulates the HEIGHT of the R wave and the shape of the QRS from beat to beat. | The R-wave amplitude rises and falls with the breathing cycle. This is a genuine physiological effect, not an electrode fault. | Breathing rate, roughly 0.2 to 0.35 Hz seated. | A few percent of R-wave amplitude, but systematic. | All amplitude features measured on single beats. IMPORTANT NUANCE FOR YOUR PROJECT: this modulation IS the ECG-derived respiration signal, and EDR is one of the two core inputs for your C6 sleep-apnoea target. So for C1, C2 and C4 it is noise to be averaged out, but for C6 IT IS SIGNAL AND MUST BE PRESERVED. A denoiser that removes it will destroy your sleep-apnoea features. | Handle by averaging over many beats (median beat) for morphology features, while keeping the unaveraged beat-to-beat amplitude series for EDR. Do not filter it away globally. | Both chips equally affected. | 100 percent | N4

## 2_Feature_x_Noise
Feature (from your features workbook) | Baseline wander | Electrode motion | Muscle / EMG | Powerline 50 Hz | Impedance drift | DC offset / saturation | Quantisation | Amplifier noise floor | Note
P-wave onset / offset (P duration, P dispersion) | H | H | H | M | L | H | M | M | The P wave is the smallest wave you must delineate, so almost everything hurts it.
P-wave area / amplitude / PTFV1 / PIFV1 | H | H | H | M | H | H | H | M | YOUR TOP C1 FEATURE. It is hit hard by six of the eight noises. Treat it as the primary thing to protect.
P-wave axis | H | H | M | L | M | H | M | L | Computed from net P area, so it inherits every baseline error.
PR interval | M | H | M | L | L | H | L | L | A timing measure, so amplitude-scaling noises mostly spare it.
QRS onset / offset (QRS duration) | M | H | H | M | L | H | M | L | EMG blurs the exact boundary; the tall QRS resists baseline drift.
QRS amplitudes (SV1, RV5, Sokolow-Lyon, Cornell) | H | H | M | L | H | H | M | L | Pure amplitude features, measured from the baseline, so drift and scaling hit them directly.
QRS axis | H | H | M | L | M | H | M | L | Net-area based, so baseline errors propagate straight into the angle.
Ventricular activation time (VAT) | L | H | M | M | L | H | M | L | Timing from QRS onset to R peak, so it survives amplitude noise well.
Fragmented QRS (fQRS) | M | H | H | H | L | H | H | M | HIGHEST FALSE-POSITIVE RISK IN THE WHOLE LIST. EMG fuzz, 50 Hz ripple and quantisation staircases all imitate the notching that defines fQRS.
J point and ST level (ST-T change) | H | H | M | M | M | H | M | L | The single worst victim of baseline wander, and also of over-aggressive high-pass filtering meant to remove it.
T-wave amplitude | H | H | M | L | H | H | M | L | Amplitude from baseline, same story as the QRS amplitudes.
T-wave end (QT, QTc) | H | H | H | M | L | H | M | M | The tangent method needs both a clean T downstroke and a trustworthy baseline to cross. Noise attacks both.
QT dispersion / QTc dispersion | H | H | H | M | L | H | M | M | ALREADY FRAGILE BEFORE ANY NOISE. It is max minus min across leads, so it amplifies the single worst lead. Expect this to break first.
TpTe, Tp-e/QT, Tp-e/QTc | H | H | H | M | L | H | M | M | Depends on both T peak and T end. Note that the RATIO forms resist impedance drift, because scaling cancels.
T-wave axis | H | H | M | L | M | H | M | L | Net-area based, same as QRS axis.
frontal QRS-T angle | H | H | M | L | M | H | M | L | YOUR TOP C4 FEATURE. It inherits errors from BOTH the QRS axis and the T axis, so its errors compound.
R-peak detection and heart rate | L | H | M | L | L | H | L | L | Robust to most things, but electrode motion creates FALSE BEATS, which is the worst failure mode because it corrupts everything downstream.
RR / NN series and HRV time domain | L | H | M | L | L | H | L | L | One false or missed beat ruins RMSSD and pNN50 for the whole window. Beat detection accuracy matters far more than waveform cleanliness here.
HRV frequency domain (LF, HF, LF/HF) | L | H | M | L | L | H | L | L | Same dependency. Note that baseline wander and the HF band overlap in frequency, but HRV is computed from beat TIMES, not the waveform, so wander mostly does not reach it.
HRV non-linear (SD1, SD2, entropy, DFA) | L | H | M | L | L | H | L | L | Entropy measures are especially sensitive to spurious beats, because a false beat looks like genuine unpredictability.
ECG-derived respiration (EDR) | S | H | M | L | M | H | M | L | SPECIAL CASE. Marked S because baseline wander is partly the SIGNAL here, not noise. Your C6 target depends on it. Do not let the denoiser remove it.
Microvolt T-wave alternans (TWA) | [8 empty] | HIT HARD BY ALL EIGHT. It is measured in single microvolts, which is below the AD8232 noise floor of about 14 microvolts peak-to-peak. Realistically not measurable on this hardware. | bar: S to L
QT variability index (QTVI) | H | H | H | M | L | H | M | M | Needs an accurate T end on EVERY beat, with no averaging available to help. Very demanding.
U wave | [8 empty] | Tiny and low-frequency, so every noise reaches it. Expect this to be unmeasurable in casual home recordings. | bar: S to L
Legend
H | S: High impact - this noise substantially corrupts this feature
M | S: Medium impact - measurable degradation
L | S: Low impact - largely survives this noise
S | S: SIGNAL, not noise - this component carries information you need (see EDR row)

## 3_Hardware_AD8232_vs_ADS1293
Specification | AD8232 (Analog Devices) | ADS1293 (Texas Instruments) | Why it matters for your project
ECG channels measured at once | 1 (single lead). The part is literally titled 'Single-Lead, Heart Rate Monitor Front End'. | 3 simultaneous channels, with Wilson and Goldberger terminal generation built in. | THIS IS THE MOST IMPORTANT ROW IN THIS WORKBOOK. See the note below the table.
Output type | Analog. You supply your own ADC. | Integrated 24-bit converters. | With the AD8232 your resolution is set by whatever ADC you attach, which is a common hidden bottleneck.
Resolution | Set by your external ADC (often 10-bit on an Arduino). | 24-bit. | Directly controls whether P-wave area and fine morphology survive digitisation.
CMRR (DC to 60 Hz) | 80 dB minimum, 86 dB typical. | Right-leg drive provided; see the TI datasheet for the figure. | Sets how well 50 Hz mains hum is rejected before it ever reaches your filter.
Input voltage noise | 100 nV per root Hz at 1 kHz. 12 uV peak-to-peak over 0.1-10 Hz. 14 uV peak-to-peak over 0.5-40 Hz. | Lower-noise design. | 14 uV p-p is about 6 to 14 percent of a typical P wave, and it is ABOVE the amplitude of T-wave alternans.
DC differential input range | Plus or minus 300 mV. | Much larger headroom from 24-bit conversion. | Electrode half-cell offsets can exceed 300 mV, which saturates the AD8232 and loses data outright.
Saturation recovery | Fast restore circuit. Feedback recovery switch on-time 110 ms. | Programmable. | Every saturation event costs you roughly 0.1 s of unusable trace on the AD8232.
Lead-off detection | AC method: 100 kHz excitation, 200 nA peak-to-peak, 10-20 Mohm threshold, 110 us detection delay. | Both AC and DC lead-off detection. | FREE GROUND TRUTH. Log this flag and you get automatic labels for your signal-quality work in step 3.
Filtering | Two-pole high-pass built in, plus an uncommitted op-amp for a three-pole low-pass. | Digital filter cutoff programmable from 5 Hz to 1280 Hz per channel, independently. | The AD8232's high-pass is tuned for heart-rate monitoring, not diagnostic ECG, so it risks distorting your ST segment.
Supply current | 170 uA typical, 230 uA maximum. | Low-power design for portable use. | Not a signal-quality issue, but relevant if you battery-power the device to avoid mains coupling.

## 4_References
Ref ID | Source | Link | Peer-review status
N1 | Moody GB, Muldrow WK, Mark RG. A noise stress test for arrhythmia detectors. Computers in Cardiology. 1984;11:381-384. Source of the three standard noise recordings (bw, em, ma). | https://physionet.org/content/nstdb/1.0.0/ | Peer-reviewed (conference)
N2 | MIT-BIH Noise Stress Test Database (NSTDB) v1.0.0, PhysioNet. Contains three half-hour real noise recordings: bw (baseline wander), em (electrode motion), ma (muscle artefact), recorded from electrodes placed so that no ECG is visible. | https://physionet.org/content/nstdb/1.0.0/ | Curated database
N3 | Friesen GM, Jannett TC, Jadallah MA, Yates SL, Quint SR, Nagle HT. A comparison of the noise sensitivity of nine QRS detection algorithms. IEEE Trans Biomed Eng. 1990;37(1):85-98. Tests five noise types: EMG, 60 Hz powerline, respiratory baseline drift, abrupt baseline shift, and a composite. | https://www.robots.ox.ac.uk/~gari/teaching/cdt/A3/readings/ECG/Friesen1990.pdf | Peer-reviewed
N4 | Zhang J, Guo Y, Dong X, Wang T, Wang J, Ma X, Wang H. Opportunities and challenges of noise interference suppression algorithms for dynamic ECG signals in wearable devices: A review. Measurement. 2025;250:117067. | https://www.sciencedirect.com/science/article/abs/pii/S0263224125004269 | Peer-reviewed
N5 | Kligfield P, Gettes LS, Bailey JJ, et al. Recommendations for the Standardization and Interpretation of the Electrocardiogram, Part I: The Electrocardiogram and Its Technology. Circulation. 2007;115(10):1306-1324. Source of the 0.05 Hz low-frequency cutoff rule. | https://www.ahajournals.org/doi/10.1161/CIRCULATIONAHA.108.191095 | Peer-reviewed
N6 | Buendia-Fuentes F, et al. Prevalence and clinical implications of improper filter settings in routine electrocardiography. Am J Cardiol. Shows that raising the high-pass cutoff distorts the ST segment; only 25 percent of routine ECGs met bandwidth standards. | https://pubmed.ncbi.nlm.nih.gov/17317378/ | Peer-reviewed
N7 | Luo S, Johnston P. A review of electrocardiogram filtering. J Electrocardiol. 2010;43(6):486-496. | https://www.sciencedirect.com/science/article/abs/pii/S0022073610002852 | Peer-reviewed
N8 | Source-level motion artifact suppression in wearable electrodes: from underlying mechanisms to advanced strategies. Review describing the three motion-artifact mechanisms: impedance instability, biopotential (half-cell) instability, and friction/triboelectric effects. NOTE: I verified the title, content and link but not the journal name. | https://www.sciencedirect.com/science/article/pii/S2590006426004850 | Review - journal not verified
N9 | AD8232 Single-Lead, Heart Rate Monitor Front End, Data Sheet Rev. A. Analog Devices. Source of all AD8232 numbers quoted in this workbook. | https://www.analog.com/en/products/ad8232.html | Manufacturer datasheet
N10 | ADS1293 Low-Power, 3-Channel, 24-Bit Analog Front-End for Biopotential Measurements. Texas Instruments. | https://www.ti.com/product/ADS1293 | Manufacturer datasheet
N11 | Dependence of Skin-Electrode Contact Impedance on Material and Skin Hydration. Sensors. 2022;22(21):8510. | https://www.mdpi.com/1424-8220/22/21/8510 | Peer-reviewed
N12 | ECG signal quality in intermittent long-term dry electrode recordings with controlled motion artifacts. Sci Rep. 2024;14. | https://www.nature.com/articles/s41598-024-56595-0 | Peer-reviewed
N13 | Fully Screen-Printed and Gentle-to-Skin Wet ECG Electrodes with Compact Wireless Readout. Reports roughly 30 percent signal degradation after 16 hours and roughly 35 percent R-peak amplitude loss after 24 hours as gel dries. | https://pmc.ncbi.nlm.nih.gov/articles/PMC11022287/ | Peer-reviewed
N14 | Preprocessing and Denoising Techniques for Electrocardiography and Magnetocardiography: A Review. | https://pmc.ncbi.nlm.nih.gov/articles/PMC11591354/ | Peer-reviewed
N15 | De Luca CJ, et al. and standard surface-EMG references: surface EMG occupies roughly 5-500 Hz with dominant energy 50-150 Hz, overlapping the diagnostic ECG band of 0.05-150 Hz and the QRS energy band. | https://www.delsys.com/downloads/TUTORIAL/semg-detection-and-recording.pdf | Technical tutorial
N16 | Investigating the Generalizability of ECG Noise Detection Across Diverse Data Sources and Noise Types. arXiv:2502.14522. | https://arxiv.org/html/2502.14522v1 | PREPRINT - not peer reviewed
N17 | Rautaharju PM, Surawicz B, Gettes LS. AHA/ACCF/HRS Recommendations, Part IV: The ST Segment, T and U Waves, and the QT Interval. Circulation. 2009;119(10):e241-e250. Defines where ST is measured, which is what baseline wander attacks. | https://pubmed.ncbi.nlm.nih.gov/19228821/ | Peer-reviewed

## 5_Read_this_first
Read this first
The single biggest practical problem with your plan | The AD8232 measures ONE lead. But most of the features in your All_features_of_ECG workbook need more than one lead, and several need specific chest leads. With a single-lead AD8232 you CANNOT compute: SV1 (needs V1), RV5 (needs V5), Sokolow-Lyon, Cornell voltage, P-wave area in V3 (your top C1 predictor), QT dispersion, QTc dispersion, P-wave dispersion, Tp-e dispersion, R-wave progression, fragmented QRS across contiguous leads, or interatrial block. You also cannot compute the frontal QRS-T angle, because that needs leads I and aVF.
What each chip actually unlocks | AD8232 (1 lead): heart rate, RR and NN series, all HRV features, ECG-derived respiration, and single-lead intervals such as QRS duration, QT, QTc, TpTe and their ratios. That is enough for your C6 sleep-apnoea target and nothing else. ADS1293 (3 channels): you can record leads I, II and III simultaneously and derive aVR, aVL and aVF from them, which DOES give you the frontal QRS-T angle and therefore your C4 preeclampsia target. It still does not give you V1, V3 or V5 unless you place chest electrodes and record in more than one session.
Recommendation | If you intend to keep all four target diseases, the ADS1293 is the only one of the two that can support C4, and neither supports C1 as specified (P-wave area in V3 needs a chest lead). Either move to a chest-lead configuration, or accept that C1 and C2 features will have to be validated on the public 12-lead datasets rather than on your own home recordings.
How the ranking was decided | Ranked by prevalence multiplied by damage, for your stated setting: seated patient, small movements, small muscle movements, electrode movements, at home. Powerline interference is extremely common but sits at rank 4 rather than higher, because a notch filter removes it reliably, whereas ranks 1 to 3 overlap the signal and cannot be filtered out cleanly.
Mains frequency | Bangladesh runs on 50 Hz. Most published ECG denoising papers use American data and assume 60 Hz. If you copy a published filter without changing this, it will not remove your interference. Set your notch to 50 Hz and expect harmonics at 100 and 150 Hz.
Honest limitations of this workbook
Assumption I made | The quantisation worked example assumes a 10-bit ADC over a 3.3 V range and a total gain of 100 or 1000. Your actual board may differ, so check it against your own hardware. The arithmetic is shown so you can redo it with your numbers.
Sources I could not fully verify | Reference N8 (motion artefact mechanisms review): I verified the title, content and link, but not the journal name, so I have not stated one. Reference N16 is an arXiv PREPRINT and has not been peer reviewed; it is included for its noise taxonomy only. Every other reference in the References sheet was checked against the published record.
Not covered here | I have not included quantitative signal-to-noise ratios for each noise type in your specific setup, because that depends on your electrodes, skin prep and room, and no published figure would transfer honestly. Measure it once you have hardware.